# Thermal Effects and Temperature Profiles in Capillary Electrophoresis

by

Christopher John Evenhuis

A thesis submitted in fulfillment of the requirements for the degree of Doctor of Philosophy

January 2007

# Declaration

To the best of my knowledge, this thesis contains no copy or paraphrase of material previously written or published, except where due reference is made.

This thesis may be made available for loan, and limited copying in accordance with the Copyright Act 1968.

Christopher John Evenhuis

10 January 2007

## Acknowledgements

I would like to thank the following people for their help and encouragement over the last three years.

My primary supervisor Dr Miroslav Macka for his guidance, enthusiasm and generosity with his time and resources to support my project.

My research supervisor Dr Rosanne M Guijt for her help, encouragement and invaluable proofreading of numerous drafts of manuscripts.

My associate supervisors Prof. Paul R. Haddad and Prof Philip J. Marriot for their wisdom and guidance and polishing of manuscripts.

Dr Cameron Johns for getting me started in capillary electrophoresis.

Dr Michael C. Breadmore for many helpful suggestions and instructing me on how to maintain and repair my capillary electrophoresis instrument.

Dr Emily F. Hilder for her assistance and encouragement.

All members of the Australian Centre for Research on Separation Science

John Davis and Paul Waller for electronic technical assistance.

Dr I. Koev for the gift of PMMA capillary.

Xiangchun Xuan and Donqing Li for supply of original data files for Figures Error: Reference source not found to Error: Reference source not found.

Vlastmil Hruška for assistance in performing ionic strength corrections to electrophoretic mobilities and for useful collaboration in Prague.

Mrs. Violet Giblin for her continued interest and enthusiasm.

My parents Dr A. J. Evenhuis and Dr E.K. Evenhuis for their encouragement.

My friend and prayer partner Omar Venegas for his encouragement and prayer support throughout the course of my research.

My daughters Ruth and Elyse for their continued interest and encouragement.

My wife Rosanne for her help in supporting our family during the course of my research and for her constant friendship and encouragement.

Finally I thank God for leading me back in to academia and for supplying all my needs.  This work is dedicated to Him.

Trust in the LORD with all your heart, and lean not on your own understanding; in all your ways acknowledge Him, and He shall direct your paths.  Proverbs 3:5-6 (NKJ)

## Glossary of Terms

| Abbreviation | Quantity | Unit |
| --- | --- | --- |
| α | Thermal diffusivity | m2s-1 |
| β | Reciprocal of film temperature | K-1 |
| γ | Temperature coefficient of electrical conductivity | ºC-1 |
| γ∗ | Temperature-dependence of the diffusion coefficient | Dimensionless |
| δt | thickness of the thermal boundary layer | m |
| ∂p∂Li | Pressure gradient in section i of the capillary. | kgm-2s-2 |
| ΔTAir | Temperature difference across air layer surrounding capillary (see Error: Reference source not found) | ºC |
| ΔTAcross Wall | Temperature difference across the capillary wall (sum of ΔTFS and ΔTPI (see Error: Reference source not found) | ºC |
| ΔTAt Axis | Change of electrolyte temperature at capillary axis as a result of Joule heating | ºC |
| ΔTCoolant | Temperature difference across coolant layer surrounding capillary | ºC |
| ΔTFS | Temperature difference across fused-silica wall (see Error: Reference source not found) | ºC |
| ΔTMean | Change in mean temperature of the electrolyte as a result of Joule heating (see Error: Reference source not found) | ºC |
| ΔTPI | Temperature difference across poly(imide) coating (see Error: Reference source not found) | ºC |
| ΔT(r) | Temperature difference between electrolyte and the inner wall of the capillary at a distance r from the axis. | ºC |
| ΔTRadial | Radial temperature difference across electrolyte (see Error: Reference source not found) | ºC |
| ΔTRadi | Radial temperature difference across electrolyte in section i of the capillary (see Error: Reference source not found) | ºC |
| ΔTref | Characteristic temperature rise | ºC |
| ΔTWall | Change of electrolyte temperature near capillary wall as a result of Joule heating | ºC |
|  | Electrical permittivity of electrolyte | Fm-1 |
| 0 | Electrical permittivity of a vacuum | 8.8542 x 10-12 Fm-1 |
| εWall | Electrical permittivity of electrolyte neat the inner wall of section i of the capillary (see Error: Reference source not found) | Fm-1 |
| r | Dielectric constant of water | Dimensionless |
| ζ | Zeta potential | V |
|  | Dynamic viscosity of electrolyte | kgm-1s-1 |
| ηiwall | Viscosity of electrolyte at the wall in section i of the capillary (see Error: Reference source not found) | kgm-1s-1 |
| κ | Electrical conductivity of electrolyte | Sm-1 |
| κ | Average electrical conductivity of electrolyte | Sm-1 |
| κ0 | Electrical conductivity of electrolyte at set temperature | Sm-1 |
| λ | Thermal conductivity of electrolyte | Wm-1K-1 |
| λAir | Thermal conductivity of air | Wm-1K-1 |
| λFS | Thermal conductivity of fused-silica | Wm-1K-1 |
| λPI | Thermal conductivity of poly(imide) | Wm-1K-1 |
| λWall | Thermal conductivity of wall | Wm-1K-1 |
| EOF | Electroosmotic mobility | m2s-1V-1 |
| µEOF0 | Electroosmotic mobility free from Joule heating | m2s-1V-1 |
| µEOF0(25 ºC) | Electroosmotic mobility free from Joule heating at 25 ºC | m2s-1V-1 |
| ep | Electrophoretic mobility | m2s-1V-1 |
| ep0 | Electrophoretic mobility at zero ionic strength (equivalent to limiting ionic mobility) | m2s-1V-1 |
| ep(0 Wm-1) | Effective electrophoretic mobility free from Joule heating effects | m2s-1V-1 |
| epA | Electrophoretic mobility of analyte ion | m2s-1V-1 |
| epB | Electrophoretic mobility of co-ion in electrolyte | m2s-1V-1 |
| µepi Wall | Electrophoretic mobility of ion at the inner wall of section i of the capillary (see Error: Reference source not found) | m2s-1V-1 |
| epR | Electrophoretic mobility of counter-ion in electrolyte | m2s-1V-1 |
| µosm | Electroosmotic mobility | m2s-1V-1 |
| ν | Kinematic viscosity | m2s-1 |
| ρ | Density | kgm-3 |
| σ | Standard deviation of Gaussian peak | s |
| σw | Charge density | Cm-2 |
| σw | Mean charge density | Cm-2 |
| ψ | Electrical potential | V |
| ωi | Radial viscosity distribution in section i of the capillary (see Error: Reference source not found) | Dimensionless |
| a | Gradient of graph of power versus increase of electrolyte temperature | WK-1 |
| a* | Reciprocal of a | KW-1 |
| A | Area | m2 |
| BiOA | Overall Biot number | Dimensionless |
| Bi'OA | Average of experimentally-determined Biot numbers calculated at 25 kV and 30 kV | Dimensionless |
| c | Molar concentration | molL-1 |
| cA | Molar concentration of analyte ion | molL-1 |
| C | Constant used for calculation of the Nusselts number | Dimensionless |
| CA | Charge transported of analyte per unit volume | CL-1 |
| CEC | Capillary electrochromatography | Analytical technique |
| CGC | Capillary gas chromatography | Analytical technique |
| dFS | External diameter of fused-silica | m |
| di | Internal diameter of capillary | m |
| do | External diameter of capillary | m |
| D | Diffusion coefficient | m2s-1 |
| DEff | Effective dispersion coefficient | m2s-1 |
| DTaylor | Taylor-Aris dispersion coefficient | m2s-1 |
| DWall | Diffusion coefficient at the wall, | m2s-1 |
| E | Electrical field strength | Vm-1 |
| Ei | Electrical field strength in section i of the capillary (see Error: Reference source not found) | Vm-1 |
| EDL | Electrical double layer | (Concept) |
| EOF | Electroosmotic flow | (Concept) |
| F | Faraday’s constant | 96487 Cmol-1 |
| FEP | Fluorinated ethylene-propylene copolymer | Capillary Material |
| FS | Fused-silica | Capillary Material |
| G | Conductance | S |
| Gr | Grashof number | Dimensionless |
| h | Heat transfer coefficient | Wm-2K-1 |
| hActive | h for the actively cooled section of the capillary | Wm-2K-1 |
| hPassive | h for the passively cooled capillary sections | Wm-2K-1 |
| H | Height of theoretical peak | m |
| HaxT | Contribution to theoretical plate height due to effects of axial temperature differences | m |
| HJoule | Contribution of Joule heating to the theoretical plate height | m |
| HTaylor | Contribution to the theoretical plate height stemming from Taylor dispersion | m |
| I | Electric current | A |
| Is | Ionic strength | molL-1 |
| k | Boltzmann constant | 1.3807 x 10-23 JK-1 |
| k1 | Defined constant k1 = G0/2LhOAri. | Dimensionless |
| k* | Autothermal parameter | Dimensionless |
| k*’ | Critical value of autothermal parameter | Dimensionless |
| kTaylor | Constant which depends on the geometry of the lumen | 1/48 ≈ 0.0208 for circular capillary |
| K | Capacity factor | Dimensionless |
| KD | Advective dispersion coefficient | m2s-1 |
| Kdiss | Dissociation constant | Not applicable |
| KE | Electromigrational dispersion coefficient | m2s-1 |
| K* | Overall dispersion coefficient | m2s-1 |
| Ldet | Length of the capillary to the detector | m |
| Ltot | Total length of the capillary | m |
| M | Hydrodynamic conductivity | m3skg-1 |
| Mi | Hydrodynamic conductivity in section i of the capillary (see Error: Reference source not found) | m3skg-1 |
| n | Constant used for calculation of the Nusselts number | Dimensionless |
| n | Refractive index | Dimensionless |
| N | Number of theoretical plates | Dimensionless |
| Nu | Nusselts number | Dimensionless |
| p | Pressure | Pa |
| P | Power | W |
| P/L | Power per unit length | Wm-1 |
| PDMS | Poly(dimethylsiloxane) | Material for chips |
| Pe | Peclet number | Dimensionless |
| PEEK | Poly(etherether ketone) | Capillary Material |
| PI | Poly(imide) | Capillary Material |
| PMMA | Poly(methyl methacrylate) | Capillary Material |
| Pr. | Prandtl number | Dimensionless |
| q | Charge | C |
| q* | Parameter in Debye-Hückel-Onsager equation | 0.586 |
| Q | Rate of heat production per unit volume | Wm-3 |
| r | Distance from the axis of the capillary | m |
| rh | Hydrodynamic radius of solvated ion |  |
| ri | Internal radius of capillary | m |
| rMean | Distance from axis to the location where T = TMean | m |
| R | Ideal gas constant | 8.3144 Jmol-1K-1 |
| Ra | Rayleigh number | Dimensionless |
| Re | Reynolds number | Dimensionless |
| tA | Time for analyte to migrate to detector | s |
| tEOF | Time for EOF marker to migrate to the detector | s |
| T | Temperature or absolute temperature | ºC or K |
| TAxis | Temperature of the electrolyte at axis of capillary (see Error: Reference source not found) | ºC |
| TCavity | Temperature measured in instrument cavity | ºC |
| TEffective | Effective temperature of the electrolyte without Joule heating | ºC |
| TFilm | Film temperature (average of temperature of the external wall of the capillary and of the ambient temperature of the coolant) | K |
| TMean | Mean temperature of the electrolyte in the capillary (see Error: Reference source not found) | ºC |
| TSet | Set temperature of air used for active-cooling of cassette (see Error: Reference source not found) | ºC |
| TWall | Temperature of the electrolyte near the inner wall of capillary (see Error: Reference source not found) | ºC |
| TWall (EOF) | Temperature of the electrolyte near the inner wall of capillary determined using electroosmotic mobility | ºC |
| TWall (G) | Temperature of the electrolyte near the inner wall of capillary determined using conductance | ºC |
| TWalli | Temperature of the electrolyte near the inner wall of section i of capillary (see Error: Reference source not found) | ºC |
| vA | Mean velocity of analyte over cross section | ms-1 |
| vCoolant | Speed of coolant flow | ms-1 |
| vep | Electrophoretic velocity | ms-1 |
| vEOF | Velocity of the electroosmotic flow | ms-1 |
| vEOFi(y) | Electroosmotic velocity profile using a dimensionless radial coordinate | ms-1 |
| vobs | Observed velocity of analyte | ms-1 |
| vparab | Average velocity of the parabolic flow | ms-1 |
| vpdi | Hydrodynamic velocity in section i of the capillary (see Error: Reference source not found) | ms-1 |
| V | Applied voltage | V |
| Vol | Volume of electrolyte in capillary | m3 |
| w1/2 | width of peak at half height | s |
| x | Thickness of stationary air layer surrounding capillary | m |
| xS | Thickness of Stern layer (see Error: Reference source not found) | m |
| y | Dimensionless radial coordinate | Dimensionless |
| z | Valency of ion | Dimensionless |
| zA | Valency of analyte ion | Dimensionless |

## List of Publications

| Type of Publication | Number | Reference |
| --- | --- | --- |
| Papers in refereed journals | 6 | [1-6] |
| Book chapter | 1 | [7] |
| Refereed posters at international meetings | 6 | [8-13] |
| Contributions to talks at international meetings | 2 | [14, 15] |

[1] 	Evenhuis, C. J., Guijt, R. M., Macka, M., Haddad, P. R., Determination of inorganic ions using microfluidic devices, Electrophoresis 2004, 25, 3602-3624.

[2]	Guijt, R. M., Evenhuis, C. J., Macka, M., Haddad, P. R., Conductivity detection for conventional and miniaturised capillary electrophoresis systems, Electrophoresis 2004, 25, 4032-4057.

[3]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R., Internal Electrolyte Temperatures in Polymer and Fused-silica Capillaries during Capillary Electrophoresis, Electrophoresis 2005, 26, 4333-4344.

[4]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R., Variation of Zeta-Potential with Temperature in Fused-silica Capillaries Used for Capillary Electrophoresis, Electrophoresis 2006, 27, 672-676.

[5]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R., Temperature Profiles and heat Dissipation in Capillary Electrophoresis, Anal. Chem. 2006, 78, 2684-2693.

[6]	Kuban, P., Evenhuis, C. J., Macka, M., Haddad, P. R., Hauser*, P. C., Comparison of different contactless conductivity detectors for the determination of small inorganic ions by capillary electrophoresis, Electroanalysis 2006, 18, 1289-1296.

[7]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R., in: Landers, J. P. (Ed.), Handbook of Capillary Electrophoresis, CRC Press, Boca Raton 2006, p. Submitted.

[8]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R., Measuring Temperatures In Capillaries Used For CE, 28th International Symposium on Capillary Chromatography & Electrophoresis, May 22-25, 2005, Las Vegas, poster no.59

[9]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R., How Cool Is My CE?- Measuring the heat transfer coefficient in CE, International Symposium on Chromatography, Copenhagen, August 21-25, 2006 poster no. Pe01.

[10]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R.,

Get it right! Measuring electrophoretic mobilities free of Joule Heating, International Symposium on Chromatography, Copenhagen, August 21-25, 2006 poster no. Pe 03.

[11]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R.,

Post-Blast Identification of Improvised Explosive Devices, International Symposium on Chromatography, Copenhagen, August 21-25, 2006 poster no. Pe 04.

[12]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R.,

A New Approach to Measuring the Variation of Zeta-Potential with Temperature Studied in Fused-silica Capillaries in Capillary Electrophoresis, International Symposium on Chromatography, Copenhagen, August 21-25, 2006 poster no. Pe 05

[13]	Evenhuis, C. J., Guijt, R. M., Macka, M., Marriott, P. J., Haddad, P. R., Quantifying temperature increases due to Joule heating in electrodriven separations, International Symposium on High Performance Liquid Phase Separations & Related Techniques - HPLC-2005,  Stockholm, June 26-30, 2005 poster no.7:25

[14]	Evenhuis, C. J., Johns, C.,Yang, W.-C., Schuman, P., Guijt, R. M., Haddad, P. R., Marriott, P. J., Macka, M., New Approaches to Simultaneous Separations of Inorganic Anions and Cations by Capillary Electrophoresis: Use of High-Magnitude EOF Capillaries and UV-transparent Fluoropolymer Capillary, 25th International Symposium on Capillary Chromatography (ISCC), Riva del Garda, 31 May - 4 June 2004, oral keynote presentation No. KNL.10.

[15]	Evenhuis, C. J., Guijt, R. M., Haddad, P. R., Marriott, P. J., Macka, M.,

At what effective temperatures do we actually run electro-driven separations? International Symposium on Separation Science, Parubice, 12-15 September 2005, oral presentation No. L33, p. 43.

Table of Contents

Thermal Effects and Temperature Profiles in Capillary Electrophoresis	i

Declaration	ii

Acknowledgements	iii

Glossary of Terms	v

List of Publications	x

Abstract	1

1.	Introduction and Literature Review:	3

1.1.	Basic Theory of Electrophoresis	3

1.2.	Electroosmotic Flow	4

1.2.1	Origin of Electroosmotic flow	5

1.3.	Heating Effects in Electrophoresis	7

1.3.1	Quantifying Temperature Differences	8

1.3.2	The heat transfer coefficient (h)	13

1.3.3	Investigations of Joule heating and cooling efficiency	17

1.4.	Joule Heating and Separation Efficiency	21

1.4.1	The Effect of Electrolyte Temperature on Separation Efficiency	22

1.4.2	The Effect of Radial Temperature Differences on Separation Efficiency	24

1.4.3	The Effect of Axial Temperature Differences on Separation Efficiency	34

1.5.	Temperature Measurement in CE	40

1.6.	Project Aims	48

2	Experimental	50

2.1	Instrumentation	50

2.2	Reagents	51

2.3	Conductance Measurements	51

2.4	Electroosmotic Mobility Measurements	51

2.5	Temperature Measurements	52

3	The Radial Temperature Profile of the Electrolyte During CE	54

3.1	Modeling the Radial Temperature Profile	54

3.2	Conclusions	57

4.	Using Conductance and Electroosmotic Mobility as Temperature Probes for the Mean Electrolyte Temperature	58

4.1	Introduction	58

4.2	Calibrating Conductance as a Temperature Probe	58

4.2.1	Experimental	58

4.2.2	Results and Discussion	60

4.2.3	Conclusions	67

4.3	Calibrating Electroosmotic Mobility as a Temperature Probe	67

4.3.1	Experimental	67

4.3.2	Results and Discussion	68

4.3.3	Conclusions	71

4.4	Comparison of µEOF and G as temperature probes	71

4.4.1	Introduction	71

4.4.2	Results and Discussion	71

4.4.3	Conclusions	74

4.5	Using G to find TWall	74

4.5.1	Introduction	74

4.5.2	Results and Discussion	75

4.5.3	Conclusions	75

4.6	Calculation of a complete temperature profile for a fused-silica capillary	75

4.6.1	Introduction	75

4.6.2	Theory	78

4.6.3	Experimental	79

4.6.4	Results and Discussion	80

4.6.5	Conclusions	83

4.7	Investigation of an alternative method for determining h.	85

4.7.1	Introduction	85

4.7.2	Theory	85

4.7.3	Experimental	86

4.7.4	Results and Discussion	87

4.7.5	Conclusions	91

5	The Variation of Zeta-Potential with Temperature	92

5.1	Introduction	92

5.2	Experimental	94

5.3	Results and Discussion	94

5.4	Conclusions	96

6	Temperature Increases in Polymer Capillaries	99

6.1	Introduction	99

6.2	Theory	101

6.3	Experimental	102

6.4	Results and Discussion	103

6.5	Conclusions	111

7	Measurement of Electrophoretic Mobilities Free from the Effects of Joule Heating	112

7.1	Introduction	112

7.2	Experimental	116

7.2.1	Chemicals	116

7.2.2	Procedures	117

7.3	Results and Discussion	117

7.4	Conclusions	122

8	General Conclusions	123

9	References:	126

10	Appendices	136

Appendix 1 - Derivation of Modified Debye-Hückel-Onsager Equation	136

Appendix 2 – Calculation of h for Passively-cooled Capillaries	139

Appendix 3 – Derivations of Eqns 4- 27 and 4- 28	141